## Sheet1
Sr.No | Item Name and Specifications Rate per Unit | Rate per Unit including all taxes | Items Manufacturer/ company name (compulsory) | Remark
101 | SKIN & VD Laser For Pigmentation, fungal nails Q-Switched Nd:YAG Laser LASER Q – Switched Nd: YAG laser Dual Wavelenght Wavelenght - 1064 nm, 532 nm Beam Delivery – Direct delivery with integrated aiming beam Pulse energy – variable , upto 1000mj Fluence – variable , preferably ( 2- 14 j/cm2) Spot size – 2.5 mm, 3.5mm, 5mm ( optional 6.5 mm) Pulse rate - variable , 1- 10 Hz
102 | Electroporation device (Needle free mesotherapy) Fully automated electronic device -Power supply 9V DC, 1 A max, external supply, -Average pulse current +/- 1 ma, 2 ma, 3 ma, -Load impedance: 0-15 kohm, -Pulse frequency: 2200 hz, -Burst time: 10 msec, -Burst frequency: 50 hz
103 | E.N.T. Medical Grade television Panel size : TFT AM LCD Active Screen Size (H x V)708.48 x 398.52 mm Aspect Ratio :16:9 Resolution:2560 x 1440 Pixel Pitch :276.75 um Viewing Angle (H,V)178°, 178° Maximum Luminance:450 cd/m²(5 points avg) Contrast Ratio : 3000:1 Color : 1.07 Billion Response Time (Tr + Tf ) 20 ms Video Inputs Signals 2 x SD/HD/3G-SDI (BNC) 2 x SOG 1 x VGA (D-sub) 1 x C-Video (BNC) 1 x S-Video (DIN) 1 x Component (RGBS, YPbPr) (5 x BNC) GPIO Port SWAP, PIP / PBP1 / PBP2 Select, Record Indicator Video Outputs Signals 2 x DVI-D,
104 | Medical Grade television Panel size : TFT AM LCD Active Screen Size (H x V)708.48 x 398.52 mm Aspect Ratio :16:9 Resolution:2560 x 1440 Pixel Pitch :276.75 um Viewing Angle (H,V)178°, 178° Maximum Luminance:450 cd/m²(5 points avg) Contrast Ratio : 3000:1 Color : 1.07 Billion Response Time (Tr + Tf ) 20 ms Video Inputs Signals 2 x SD/HD/3G-SDI (BNC) 2 x SOG 1 x VGA (D-sub) 1 x C-Video (BNC) 1 x S-Video (DIN) 1 x Component (RGBS, YPbPr) (5 x BNC) GPIO Port SWAP, PIP / PBP1 / PBP2 Select, Record Indicator Video Outputs Signals 2 x DVI-D,
105 | Binocular loupes with Headband & headlight As Above (Sr. No. 3)
106 | Portable microscope Built in LED light source with an intensity greater than 55,000 lux Objective fine focusing of a range of 10 mm Microscope head can be moved any direction / position 0.6 x, 1.0 x, 1.6 x magnification Weight:<10kg Microscope should be free from locking norms Table mount property
107 | CCTV camera with display trolley for demonstrating microscopic ear surgery & endoscopic sinus surgery High definition camera with LED monitor display
108 | Portable endoscopy unit with monitor and 0°endoscope TV system :PAL Image Sensor :1/3” CCD image sensor CCD total pixels :795 (H) * 596 (V) Scanning system :625 lines, 50 fields/sec Minimum illumination :0.8 Lux Resolution :470 TVL/ 580 TVL (Enhanced) White balance :Single press Zoom function :ON/OFF Electronic shutter(A.E.S) : 1/50- 1/120,000 sec Gain control :AGC on/off Video output :Composite/ USB Source :200-220 V/ 50Hz 0° endoscope : 2.7 mm/ 4.0 mm, 18 cm, wide angled
109 | Blakesley nasal straight forcep 10cm shaft,19cm overall length, Size 0,1,2
110 | Antrum grasping forceps 360°rotable,Retrograde opening jaws 120 °, 10 cm working length
111 | Gorney suction elevator with stylet 3mm tip,19 cm
112 | Tita-prosthesis piston type loop for stapedioplasty microscopic ear surgery Functional length 7.00mm to 3.50mm Malleable Trimmable by osseous cutter
113 | Super titanium vario piston type loopmicroscopic ear surgery Self crimping effect
114 | Osseous cutter to trim osseous surgical implants used in microscopic ear surgery Having 3 parts (1) Cutting block with statim for easy handling (2) Lever having sharp edge for cutting prosthesis (3) Shortening blade (a)consisting of various length steps from 1.500mm to 5.00mm (hole length) (b) 2 holes on each step
115 | Stapes prosthesis handling forceps microscopic ear surgery With facility of catching the prosthesis hook forceps jaws having grooves -stainless steel
116 | Stapes prosthesis crimping forceps used in microscopic ear surgery Stainless steel
117 | SURGERY X Ray view box Dimension – 150*60*15 cms, Power consumption – 25-100 volt, Frequency – 50 Hz, Power input – 240 V, Net weight – 10kg, No UV emission,
118 | Proctoscope 19mm Internal Diameter & 190mm Length With Obturator, Material - Stainless steel, Weight – 80-180 grams,
119 | Pedestal light Portable single dome chargeable light with dome size 500 mm with focal diameter 15-20 cm with intensity of 1 lakh unit., Height – 200 cm.
120 | Electrosurgical cautery unit Electrocautery for surgery with all attachment, Monopolar + Bipolar + argon outlet, Argon enhanced cut and coagulation, Max power output 400 watts, Defibrillator proof output, Weight less than 4.5 kg - portable, With footpad,
121 | Suction Machine Suction machine Foot pad operated (ISI Marked, Ready to use with installation & Replace piece Three years Warranty with free service within 24 ours at our premises) CE Certified,2 zars
122 | Kidney tray Steinless Steel 8" 10” ISI mark
123 | Bowl Steinless Steel large
124 | Mosquito artery forceps Mosquito forceps cvd length 14cm Country origin should mention on each & every instrument, lot number & article number should be mention each and every instruments , it should have European CE certificate, US- FDA it should be made of German Stainless steel.
125 | Bard parker Knife handle no.3 Knife handle stainless steel 316, Jewellers forcep 8" ISI CE certified
126 | Surgical Blade no.11, 15, 22 Surgical BP blade no. 11, 15, 22Steinless steel
127 | Sponge holder Steinless steel large sponge holder, FOERSTER-BALLENGER SPONGE HOLDING FORCEPS 18 CM STRAIGHT ABDOMINAL & GENERAL
128 | Tissue cutting scissor Mayos S.S 316 tissue cutting dissection curved scissors with antiglare, with golden plated handle 8 inch CE,, ISI certified Non Corrosive
129 | Thread cutting scissor Heath suture cutting scissor, curved scissors with antiglare, stout with golden plated handle 8 inch CE,, ISI certified
130 | Needle holder Needle holder S.S316 antiglare, with golden handle,with titanium carbide grip medium, 8 inch, cE certified
131 | Babcock’s forceps TISSUE AND ORGAN HOLDING FORCEPS - BABCOCK (SMALL MOUTH DELICATE) 15CM / 6"
132 | Toothed forceps Tooth Forceps , anti glare, 8 inch/6inch/10 inch CE certified
133 | Plain forceps Toothless Forceps Plain 8" CE CERTIFIED
134 | Allis forceps Allis forceps 6 inch, Straight steinless steel, CE certified
135 | Medium sized artery forceps DISCETING ARTERY FORCEPS (CURVED MEDIUM) CE MARK germany US FDA APPROVED
136 | Autoclave Autoclave Machine Double Drum SS Electrically operated with stand I.S.I. No. 6" x 11", 1 Nos., Steinless steel, Highest working temperature 129’c, Pressure gauge indicator, Dual heating system, Dimension of sterilizer chamber – 12”*12”
137 | BP Machine Sphygmomanometer I.S.I. Without stand, Mercurial, Weight – 399 grams
138 | Stethoscope Stethoscope I.S.I.
139 | Weight machine 30*30 cm size, Display size 6*3 cm, Carrying capacity – 150 kg
140 | Height scale Length – 2000 mm, Size – 250 mm2, Weight – 200 kg, Fiber surface with 300*300 mm platform
141 | Camera for student demonstration for OT Maximum flexibility in minimum freely moveable space with sterile handle & wireless HD video transmission camera support for 2 screens.
142 | Operation theatre ceiling light Fully faceted multiparabolic refractor with 3360 facets and illumination intensity up to 160 kLux, Light field diameter - 155*270 mm, Length of working column – 7000 mm, Sterile central handle with soft touch key panel.
143 | Pulse oximeter 195 gram weight, 4 inch display mode, 126*60*30 mm size, 10 hour battery life,
144 | Anesthesia machine Should be portable stainless steel, with large antistatic sturdy castor wheels fitted, with brakes, Gas cylinder (pin indexed) yokes with sliding stainless steel/sturdy clamping bars, for easy handling, Two Pin index yokes for connecting cylinders each for O2, N2O, Regulator two each for O2 and N2O with output pressure 4.22kg/cm2. N2O, regulator is activated only when minimum oxygen on flow, Should have pressure gauge for all gas inlets including central lines mounted on, the front panel for easy visibility, Should have audible alarm for O2 failure, N2O supply should cut off if O2 supply fails (Anti-hypoxic guard), Oxygen and Nitrous oxide should be linked either mechanically or pneumatically, to ensure a minimum of 25% oxygen delivery at all times to avoid delivery of, hypoxic mixture, Should have dual cascade type flow meter for O2 and N2O calibrated in multiple, scale, Provision to mount any two selectable vaporizer with interlocking facility to allow, use of only one vaporizer at a time, Iso-flurane vaporizer of newer generation having specifications equivalent to tech, 7 type to be provided, Non-return cum pressure relief valve when pressure exceeds 120cm of H2O, Should have change over from open circuit to closed circuit and vice versa, Should provide with oxygen flush switch, Circle absorber with corrugated reusable breathing circuit for closed circuit system, with each unit, Should have low flow anesthesia technique, Should have safety certificate from a competent authority CE / FDA (US), STQC CB certificate / STQC S certificate or valid detailed electrical and, functional safety test report from ERTL/ ISI (IS.13485). Copy of the certificate, test report shall be produced along with the technical bid - (a) Should have a provision for mount monitors on top of the machine, (b) The table top made up of stainless steel/ chemical resistant fiber, (c) Standard bains circuit : 2 nos. with each unit, (d) Reservoir bag (2liters): 3 nos. with each machine, (e) Connectors for bains circuit: 5 nos with each machine, (f) AMBU bag: 1 no. with each machine, (g) Pressure regulated valve with 5 meter hose and connector (conversion kit), for oxygen should be provided with each machine, (h) Should be supplied with driver gas hoses with necessary attachments, (colour coded),
145 | Stardust NH8 STR W/LK MIC European CE certificate, US- FDA approved it should be made of German Stainless steel
146 | Debakey BD CLP RING HDL S-CVD 5-1/8 European CE certificate, US- FDA approved it should be made of German Stainless steel
147 | MIC CLP 4-1/2 ANG 30mm jaw TIT-MTO European CE certificate, US- FDA approved it should be made of German Stainless steel
148 | COOLEY VASC CLP 6-1/2 90D SML European CE certificate, US- FDA approved it should be made of German Stainless steel
149 | Stardust MIC ring FCPS 7-1/4 STR European CE certificate, US- FDA approved it should be made of German Stainless steel
150 | Stardust MIC SUT FCPS 7-1/4 STR W/TP European CE certificate, US- FDA approved it should be made of German Stainless steel