## Sheet1
Sr.No | Item Name and Specifications Rate per Unit | Rate per Unit including all taxes | Items Manufacturer/ company name (compulsory) | Remark
351 | ORTHOPEDICS Movie Camera for demonstration of live operations
352 | Plaster tables
353 | Electric plaster cutter
354 | Basic instrumentation set for fracture.
355 | External Fixator
356 | C-Arm (Image Intensifier)
357 | Portable X-ray Machine
358 | Slide Projector
359 | Over-head projector
360 | Operation room equipment (sets)
361 | ANESTHESIOLOGY Patient examination table - specification Overall Approx size: 187 L x 51 W x 81 H. cms. Adjustable backrest. Two section Foam padded. Provided with three Drawers. Powder Coated.
362 | Blood Pressure Instrument - (Non-invasive)- Oscillometric. , Digital display , Systolic BP. , Diastolic BP. , Mean BP. , Selectable manual/ automatic mode. , Extar large cuff. , Adult cuff , Pediatric cuff. , Neonate cuff. , Cable.
363 | Height Measurement Scale Measuring range: 20 - 205 cm,Graduation: 1 mm. ,Stable. , Easy to fix.
364 | Weighing machine Spring type, Non digital, Easy to calibrate, Range of 10- kg to 150 kg
365 | Multiparameter Monitors with Capnography (1) Conforming to International standards US FDA and CE. Certificates To be attached . [CRITICAL] (2) IS0-13485- "Medical devices- - Quality management systems-- Requirements for regulatory purposes. It specifies requirements for a quality management system where an organization needs to demonstrate its ability to provide medical devices and related services that consistently meet customer requirements and regulatory requirements applicable to medical devices and related services . Certificates To be attached. [CRITICAL] (3) IEC Safety Standard Certified . To be attached. [CRITICAL] (4) Portable. [CRITICAL] (5) Six Channel [CRITICAL] (6) High Resolution Colour TFT Display with 12"or More Screen size . [CRITICAL] (7) Should work on Mains as well as battery (backup- 4-8 hrs) [CRITICAL] (8) Should be able to monitor [CRITICAL] ECG , Sp02- waveform and numerics, Capnography- waveform and numerics, NIBP, Temperature, Respiration, (9) Should be able to monitor these parameters in adult, paediatric and neonatal patients. [CRITICAL] (10) ECG monitoring: [CRITICAL] 3 ECG leads permit display of lead I, II, Ill. , 5 ECG leads permit display of lead I, II, Ill, avR, avl, avF, Cascade waveform facility, ST segment analysis should be possible for one lead, Arrhythmia detection, Defibrillator and cautery protection should be provided [CRITICAL], Heart Rate : 20-300 bpm, Sweep speed- 12.5 mm/s; 25 mm/sec; 50 mm/ sec . (11) Pulse Oximetr [CRITICAL] Display of plethysmograph, Display of perfusion level indicator, Display of Sp02 values, Sp02 Range 40-100%, Pitch variation with change in oxygen saturation, Pulse Rate range 30-200 bpm, Accuracy to be mentioned, Type of probe: NELCOR original OR MASIMO original (pl. specify) (with certificate issued by principal manufacturer of the probe for the quoted model ), (12) Capnography : [CRITICAL] Waveform display, Digital display for EtC02, FiC02 & RR, Should be micro stream technology (which is a variant of side stream type), Flow rate requirement should be Appro x imate 50 -60 ml per Minute (flow range to be mentioned ), ETC02 range should be 0-100 mm of Hg, FiCo2 range should be 0-20 mm of Hg, Carbon dioxide Values in mm of Hg, torr, kPa units (13) Non Invasive Blood Pressure {NIBP) [CRITICAL] Oscillometric method for adult,child& neonate, Measurement and display of systolic, diastolic and mean pressure values of NIBP measurement, Range 30-220 mm of Hg, User selectable alarm settings, 3 Modes of operation: Manual, STAT (continuous Sminute operation) , automatic (selectable time interval 2- 90minutes), Accuracy to be mentioned (14) TemQerature monitoring: [CRITICAL] Dual, Units (Degree C and Degree F) selectable, Display ofTl, T2,and differential temperature should be user selectable, Temp . Range 0-50 Deg. C, Probes should be YSI series, Reusable Skin temperature probe should be available, Reusable Nasopharyngeal temperature probe should be available, (15) Respiration: [CRITICAL] RR range : 2-60 bpm, Display of respiration waveform with respiratory rate, Sourced through ECG cable or C02, Priority to C02, Apnea alarms should be provided. (16) lnbuilt recorder [CRITICAL] (17) ALARMS : [CRITICAL] Should have audio visual alarms, Should have Alarm recall facility, Auto setting of alarm limits, Alarm storage (18) Trends and storage : [CRITICAL] 48 Hours/100 patients, Graphical & Tabular Trends, NIBP- Last 750 reading tabular form, Alarm Recall : Last 20 Alarms events with details (19) Upgradable for Invasive BP 1, Invasive BP 2. An undertaking will be given .Rates to be given in commercial bid. [CRITICAL] - Upgradable for cardiac output monitoring, anaesthetic agent gas monitoring. An undertaking will be given[Non-CRITICAL](20) Demo mode for training the staff. [CRITICAL] (21) Operating manual. [CRITICAL] (22) lnbuilt voltage stabilizer [NON-CRITICAL] (23) Standard supply : (Should be supplied with-) [CRITICAL] 3 lead ECG cable - 1, 5 lead ECG cable - 1, Reusable Sp02 Adult probe with extension cable- 1, Reusable pediatric Sp02 probe - 1, Reusable neonatal Sp02 probe - 1, Connectors for capnography monitoring, for nasal & ET Type, fo r Adult & Pediatric - 1each, NIBP cuff with one hose (1for Adult), NIBP cuff (1for Child), NIBP cuff (1for Neonatal), Reusable Skin temperature probe (YSI Series)- 1, Reusable Nasopharyngeal temperature probe (YSI Series)- 1, Reusable.Rectal temperature probe (YSI Series) - 1, Power cable - 1, Operating manual. - 1, Service manual – 1, Operation manual with user demonstration video CD.
366 | Ultrasound machine (1) Latest Technology all - Digital Hand held Colour Doppler System suitable for Central and peripheral line placements, Epidurals, Pain Management, Nerve Blocks and Tran esophageal echocardiography scanning applications. (2) The unit should be compact, lightweight and portable. Specify weight and dimensions. Weight should not exceed 3.5 kg excluding cart and accessories. (3) Should have Broad Band Frequency Transducer Technology. (4) Should have B Mode, M-Mode, Colour, Angio, Pulsed Wave Doppler Mode and Continuous Wave Doppler Mode of Imaging. (5) Should have Tissue Harmonic Imaging facility. (6) Should have fast boot up less than 30 seconds for critical situation. (7) System with inbuilt rechargeable battery will be preferred and longer battery life. Should have AC Mains adaptor / battery charger. (8) It should have inbuilt foldable alphanumeric keyboard, can be sanitize to tack in to OT. (9) Flat LCD/ TFT monitor of at least 10 inches (10) Should onboard storage of at least 10000 images. Storage in BMP and AVI format should be possible. Sorting of database with patient name and date should be possible. And have facility to transfer images to computer. (11) Should have cine memory and clip storage facility. (12) Should DICOM ready system (13) Video in/out, VCR or DVD for record possible, B/W Thermal Printer, output to external display. (14) 13-6 MHz Broad Band Linear Probe for Vascular access, Vascular imaging, Small parts, Regional nerve blocks (15) The price must be quoted for unit and transducers covered with comprehensive onsite warranty for five years.
367 | White screen for LCD projection Size: 4 X 5 fit. Good quality: With stand
368 | Pulmonary function test (PFT ) machine 1. SIZE & WT.: Wt. < 3 kg. Potable. , 10” X 5”X3” maximum. 2. PARAMETERS MEASURED: Expiratory FVC,FEV,FEV,FVC,PEFR,FEF 3. INSPIRATORY , FIVC, PIFR. 4. POWER - Power derived from USB port 5. ACCURACY , <± 3% or 100 ml. whichever is greater. 6. VOLUME RANGE , ± 16L 7. FLOW RANGE , ± 16L /sec. 8. MEMORY - Minimum 100 cases. 9. SCREEN - LCD Screen.
369 | The large full graphic LC-display High contrast and a wide viewing angle. Display of the following information: - B165Stimulus amplitude in mA (large digits) - Stimulus duration in ms - Stimulus frequency in Hz - Load impedance in kΩ - Charge delivered (in nC, if activated) Electrode cable for all needles, lngth-100 cms or more Device for percutaneous nerve mapping to help pre-assess the puncture site. Remote Control for sterile one-hand operation Battery operated. Battery cover Test resistance Knob for sterile handling Operation Manual Carrying case
370 | Anodyne Machine 4 Flexible Therapy Pads 1 Control Unit 60 infrared diodes (890 nm) per Therapy Pad Energy Setting of 840 mw per Therapy Pad Lower Extremity Application System - if applicable Lightweight and portable Durable carrying cas Personalized telephone training 2 year warranty
371 | Anaesthesia Workstation 1. General Requirement a) Compact and modular, 3 gas Anaesthesia workstation with an integrated ventilator for adult to infants and integrated airway monitor for airway pressures and volume. b) The machine should be suitable for low and minimal flow anesthesia application with compliance compensation of breathing ckt, fresh gas flow compensation/ decoupling. c) The machine should have 2 drawers d) The anaesthesia machine, inbuilt ventilator e) vaporizer supplied should be FDA approved/ European CE marked f) The system should have upto 2 Hrs. battery backup g) System should confirm to US FDA / European CE and EN 60601-2-13 (Requirement for safety and essential performance of anaesthesia system) 2. Gas delivery system a) Should have pin index yokes for Oxygen & Nitrous Oxide besides separate connection for Central gas supply for Oxygen, b) Nitrous Oxide c) The system should be suitable to use at minimal flow upto 700ml fresh gas setting. d) Automatic cutoff of N2O by Oxygen pressure failure. e) Hypoxic guard for linear regulation of minimum oxygen concentration at 23% volume f) Audible visual oxygen failure alarm. g) Emergency Oxygen flush at 30 – 70 L/min bypassing the vaporizer. 4. Flow meter - Dual Cascade type flow meter tubes for Oxygen & N2O.Range 20 ml / min to 10 Lit/min. Calibrated in multiple scales. Single tube for air 100 ml to 14 L/ min. 5. Vaporizer - Machine should have quick mount type vaporizer for easy interchangeability, and safety with interlock facility. - Should be provided with a Temperature / pressure compensated and flow independent - Should be capable of connecting Vaporiser.for Isoflourane / Sevoflourane / Halothene / Desflorane - Vaporizer should have extended delivery range from 0 to 6 Vol. % - The vaporiser should be lifelong calibration free along with manufacturer certificate for the same 6. Breathing System - Should have fresh gas de-coupled semi closed circle absorber system. - Should have adjustable pressure relief valve from 5 to 75 mbar. - The system should have leak and compliance test (including patient hoses upto the Y piece). - Should have compact breathing system with approx 1.7 Ltr. Volume capacity. - Should have an external fresh gas outlet for connecting Magill or Bain’s circuit - The device should have port for anaesthesia gas scavenging system. 7. Anaesthesis Ventilator - The system should have inbuilt ventilator with electronically controlled and pneumatic or Piston driven technology. - Should not require changing of bellows for adult & infants. - Should have minimum screen size of 6”. - Modes: Manual/Spont, Volume Controlled - The same ventilator should be capable to be upgrade to pressure support. - Tidal Volume - PEEP : 0 ~ 20 mbar - Breathing Frequency : 4 to 60 BPM - I:E Ratio : 4:1 to 1:4 - Inspiratory pause : 0 – 50% of Ti 8. Airway monitoring - Integrated monitor for electronic monitoring and display of following parameters : Expiratory Tidal Volume Expiratory Minute volume PEEP, Peak & Mean and Plateau airway pressure Frequency Waveform display for Airway pressure. 9. Alarm limits & alarms - Adjustable high / low limits with audio and visual alarms for the following : Minute volume, Airway pressure (incl stenosis and disconnect), Insp oxygen concentration, Audio power supply fail alarm, Fail to cycle warning. 10. Machine should have RS 232 connectivity port Scope of supply - 3 gas Anaesthesia machine with Patient Monitor - Trolley with min 2 drawers - Writing surface - Pin Index yokes for O2 & N2O - Pipeline connections for all three gases - Electrically operated piston based Anesthesia ventilator - Isoflourane Vaporiser - Semiclosed breathing system - Adult & Pediatric autoclavable patient tubing’s - Anesthetic mask size – Adult & child - Central gas supply hoses (Color coded) Price to be quote for: 1. System with 1 year warranty and CMC charges for next 7 years 2. System with 5 year warranty and CMC charges for next 2 years Battery has to be covered Included for warranty period and continued CMC period Price list of all proprietary consumables applicable if any.
372 | Crash Cart 1. Overall size shall be more than 900mm L x 500mm W x 1500mm H. 2. The crash cart should be made of 25.4mmx18G Stainless steel tubular frame work. 3. Shall have Epoxy / Anti-Microbial powder paint inside and out 4. Should have dual push handles on either side 5. Should have S.S. shelves, six colored removable bins & two polystyrene lockable storage units with three drawers each 6. Facility to carry ECG Monitors, Defibrillators etc on open areas at top centre and bottom shelves 7. Should have Stainless steel saline rod fixed with. 8. Two accessory mounting brackets to mount accessories anywhere without the need of pre-threaded holes having locking arrangement. 9. Crash cart should be mounted on 12.5 cms dia non-rusting swiveling castor wheels. Two 10. Oxygen cylinder stand epoxy powder coated, on one side
373 | Monitor for vital signs General · 3.5″ color TFT display · Suitable for adult, pediatric and neonatal patient · Measurement parameters: Pulse Rate/SpO2, CO2 · Audible and visual alarms with adjustable alarm ranges · Networkable with central monitoring system · Product brochure CO2 · Technology: Infrared absorption method · Mode of Sampling: Mainstream/Sidestream · Measuring parameters: Respiration Rate, EtCO2 and InsCO2 · Typical measurement range: o EtCO2: 0~150mmHg o InsCO2: 3~50mmHg · Respiration Rate: 0~150bpm (Mainstream) · CO2 Accuracy: o 0~40mmHg ±2mmHg o 41~70mmHg ±5% of reading o 71~100mmHg ±8% of reading o 101~150mmHg ±10% of reading · CO2 Resolution: 0.1mmHg (0~69mmHg);0.25mmHg (70~150mmHg) SpO2 · SpO2 measurement range: 35%~99% · SpO2 measurement accuracy: o 75% - 99% ±2% o 50% - 74% ±3% · Low perfusion performance: 0.4%~5% · Pulse rate measurement range: 30bpm~250bpm · Pulse rate measurement accuracy: ±2bpm or 2% whichever is greater · Performance under low perfusion condition: The measurement accuracy still keeps the above specification while the blood perfusion index is as low as 0.4%. Display range of Perfusion Index: 0.2%~20% Printer (optional) · Type: Built-in thermal printer · Recording speed: 25mm/s Power Supply · Power supply: 100~250VAC, 50/60 Hz · Internal power supply: 12VDC (rechargeable) · Battery specification: 12V 2.3AH (sealed lead-acid battery)
374 | Peripheral Nerve Stimulator 1. Should have a current range from 0 .1 to 5mA. 2. Should have a resolution of 0.01mA 3. Should have a digital display for current. 4. Should have short stimulus pulse duration of 0.1ms. 5. Should have a frequency range from 0.1 to 99 Hz. 6. Should be able to monitor single twitch, Train of four, PTC and DBC. 7. Should have integrated electrode cable with lead. 8. Should be battery operated. 9. Should be supplied with 5 sets of insulated needles in varied length. 10. Should have safety certificate from a competent authority CE / FDA (US) / STQC / CB certificate / STQC S certificate or valid detailed electrical and functional safety test report from ERTL. Copy of the certificate / test report shall be produced along with the technical bid.
375 | RADIOLOGY
376 | RESPIRATORY MEDICINE Rigid Bronchoscope Technical Specification of Equipment: 1. Zero degree straight forward viewing telescope with integrated fiberoptic light transmission, diameter 4.5 mm ,working length of 50 cm length- 1 No. 2. Zero degree straight forward viewing telescope with integrated fiberoptic light transmission, diameter 2.8 mm, working length of 44 cm - 1 No. 3. Tracheoscope tube size 6.5 mm, length at least 30 cm- 1 No. 4. Tracheoscope tube size 12 mm, length at least 30 cm -1 No. 5. Tracheoscope tube size 14 mm, length at least 30 cm for application of stents -1 No. 6. Bronchoscope tube size 6.5 mm, length at least 40 cm - 1 No. 7. Bronchoscope tube size 11 mm for application of stents, length at least 40 cm- 1 No. 8. Bronchoscope tube size 12 mm, length at least 40 cm-1 No. 9. Bronchoscope tube size 14 mm, length at least 40 cm for application of stents – 1 No. 10. Optical forceps, alligator -1 No. 11. Optical forceps, cupped jaws for biopsy-2 Nos. 12. Optical forceps, universal-1 No. 13. Optical forceps for removal of coins and flat foreign body-1 No. 14. Manual forceps alligator, diameter 2.5 mm at least 50 cm length-2 Nos. 15. Manual forceps round cupped jaws for biopsy, diameter 2.5 mm at least 50 cm length-1 No 16. Manual forceps universal, diameter 2.5 mm at least 50 cm length-1 No. 17. Manual forceps for peanuts and soft foreign bodies at least 50 cm length-1 No. 18. Foreign body basket with handle > 50 cm length- I No. 19. Sponge holder forceps-1 No. 20. Cotton applicator forceps -1 No. 21. Insulated coagulation tube with connector for unipolar coagulation 1 No. 22. TONN Stent applicator system for deployment of silicon stents of diameter 14-20 mm and 11- 13 mm consisting of Folding System, Clamping Rod, Loading Rod, Introducer Tube length 42 cm, with 2-ring handle and Pusher-1 Set (Red 1 and Green 1) 23. Dedicated forceps for opening and deployment of silicon stents- 1 No. 24. One set of boogies. 25. Silicone stent (3 each) : Tracheal stent with wall thickness 1.5 mm, Thin tracheal stent with wall thickness 1.0 mm, Bronchial stent with wall thickness 1.0 mm, Carina stent with wall thickness 1.0 mm, Total carina stent with wall thickness 1.0/1.3 mm 26. Controlled Radial Expansion Balloon dilator (Wire Guided) size of 6-8 mm, 8-10mm, 10-12mm, 12-14mm-2 each. 27. Each Balloon Inflation System for CRE Balloon-1 Set 28. Should be accompanied with all accessories essential for the functioning of the equipment including: 29. Prismatic Light Deflector - 5 pieces 30. Rubber telescope guide - 20 pieces 31. Glass window plug - 10 pieces 32. Movable adaptor with sealing cap - 5 pieces 33. Injection cannula for positive pressure ventilation - 5 pieces 34. Adaptor for respirator - 10 pieces 35. Instrument guide - 10 pieces 36. Rubber tipped suction catheter (4 mm) of at least 50 cm length with adaptor- 10 37. Adjustable head rest for positioning of patients head during the procedure -1 38. Flexible suction catheters - at least 100 39. Cleaning brushes - 5 40. Equipment carrying case (for the tubes and forceps) – 1 41. Anti fog solution -30 ml - 15 bottles 42. High resolution 3 chip video camera CCU and head - 1 No. [3x1/2" CCD image sensor chip, Resolution : 750 lines - 800 lines horizontal. Picture element = 752 (h) x 582(v) pixels per chip, Min. sensitivity : 3 lux (f=1.4), AGC: + 18 db, signal tonoise ration > 60 db., Camera should have integrated parfocal zoom lens, F=25-50mm, It should have DV output and image processing module and it should have image freeze function, Programmable control buttons on camera head for controlling, Gain , white balance shutter speed, video printer and recorder, Keyboard input . for data entry through built in character, Generator, Camera should be compatible with FBAS, S-VHS and RGB, Manual or automatic exposure control (1/50 sec. —1/10000), Should have automatic white balance with storage functions, For two white balance values, Unit should be certified to IEC 601-1, CE according to MDD.] 43. LED (175-300 W) light source with cable compatible with the unit -1 No. 44. High resolution medical grade flat screen 20 inch monitor for the system -1 No. Technical Specification of Equipment: 1.The tracheal and bronchoscope tubes should be made of high quality stainless steel. 2. The assembly should include a HD (High Definition) / High quality (3 chip CCD video) camera head fully compatible with the viewing telescope. 3. The video processor provided should be compatible with the camera head and provide a high resolution output to medical grade flat screen high definition/resolution 20 inch or greater sized video monitor. 4. There tracheoscope and bronchoscope tubes should be without a distal fiber optic light carrier. 5. The tracheobronchoscope tubes should be of use with proximally insertable telescopes. 6. The bronchoscopes should be durable and should be able to be cleaned with commonly used sterilizing solutions without affecting the surface of the scope. 7. All the accessories should be compatible with the sheath 8. The complete system should be covered under warranty as per the Government Rules. 9. A dedicated imported trolley for carrying the entire system and recording equipment should be provided 10. The equipment should be USFDA or European CE approved 11. A recording system to be provided which should have facility for recoding and storage of media in both image and video format and allow transfer on removable storage (either CD or USB flash drive) 12. All metallic instruments and accessories should be autoclavable 13. Operating voltage - Power 220 V 50 Hz AC . 14. The system should include all the other possible accessories, UPS, power cables, fiber optic cables connectors etc to make the unit fully functional. 15. The system should be provided with a laptop PC (Windows 7 OS, 750 GB Hard disc drive, 8 GB RAM, Core i-7 processor) 16. Price of all the accessories as mentioned should be included within the quoted price 17. Rate list of all possible spares, accessories and consumables should be provided as part
377 | Non Invasive mechanical Ventilator 1. Non Invasive ventilator with following modes: CPAP, S, S/T,PC and T (Critical) 2. Should have IPAP: 4-30 cmH2o,EPAP: 4-25 cmH2o, CPAP: 4-20 CmH2o,Rise time: 100ms to 600 ms, Ti:0.5 to 3 sec(Critical) 3. Breath rate: Should be able to set up to 30 BPM, (Critical) 4. Should be provided with integrated heated humidifier (Critical) 5. Should have a carry bag. (Critical) 6. Compliance data should be stored on external memory card which can be accessed through non proprietary card readers ( Critical) 7. Should provide flow – based pressure relief at transition from exhalation to inhalation , transition from Inhalation to exhalation and during exhalation ( Critical) 8. System should be able to compensate for variable resistance related to different mask (Critical) 9. System should have advanced technology to prevent rain out (Critical) 10. Humidifier chamber should be isolated from the inner parts of the device to prevent damage(Critical) 11. Should be European CE or US FDA Approved (Critical) 12. Device should ensure the delivery of targeted tidal volume for noninvasive patients upto 1500 ml ( Critical) 13. Should have following Alarms: Patient Disconnect , Apnea, Low Minute Ventilation, Low Tidal Volume (Critical) 14. Automatic triggering and cycling to ensure best patient to machine synchronization ( Critical) 15. Should be provided with latex free full face reusable mask Qty 2 (Critical) 16. Should be provide with Hose Pipe Qty 2 (Critical) 17. Facility for downloading compliance data from the equipment at the hospital has to be provided (Critical) 18. Should have facilities to monitor parameters like: Respiratory rate, Exhaled tidal volume, Leak, Minute ventilation ( Critical) 19. Should have Blower and valve Technology (Critical) Operation manual with user demonstration video CD.
378 | PEDIATRICS BP Cuff Manual mercury BP Cuff Neonatal-1 Pediatric-1
379 | Pulse Oxy Meter spO2 display, pulse rate display, pulse waveform display, alarm, rechargeable battery, Pediatric and Neonatal probe
380 | Infusion Syringe Pump Battery rechargeable, LCD /TFT display, syringe range- 20-50 ml, flow range 0.1-1000ml /hr. Alarm for occlusion, syringe empty.
381 | Multipara Machine LED Monitor display 12.1 with spO2, NIBP, ECG, HR, RR, visual and audio alarm, keypad and touch screen function Rechargeable battery
382 | Nebulizer machine Piston compressor nebulizer, power-220-240v, medication capacity 5ml
383 | ECG Machine (Electrocardio graph) Portable 3 channel ECG recording, automatic and manual recording modes, power supply mains and battery operated, thermal printing on 60 mm paper, rhythm lead trace up to 100 seconds
384 | Bone Marrow Aspiration Needle Stainless steel, length 50mm size 14G, 16G, with stopper
385 | Liver biopsy Needle Stainless steel WILMS SILVERMANN / MENGHINI
386 | Kidney Tray Stainless steel 8’’ size
387 | Autoclavable Drager savina Venti. Circuit Auto clavable
388 | Venesection Tray set Scissors,Mosquito forceps, curved haemostat, scalpel no10 &11,iv catheter 14G Prolene suture 4.0,silk tie 3.0 for Pediatric patients and infants
389 | Venturi mask Paediatric venturi mask with 6 diluters, made of non-toxic PVC, six colour coded diluters for 6 concentration 24%,28%,31%35%,40%,50% with adjustable nose clip, lock& adapter for high humidity with 2m anti crush tube
390 | Non Rebreathing Mask Adult, paediatric
391 | Laryngoscope Laryngoscope different handle size, blade no- 00, 0,1, miller type STRAIGHT BLADE, LED Bulb Blade no 1,2,3,4 Macintosh type, CURVED BLADE, LED Bulb
392 | Jackson Rees Circuit 18 inches corrugated tube of 15mm diameter, elbow Washington T piece for connecting fresh gas tube, 22 mm plug in mount to connect Boyles apparatus, rebreathing bag 0.5 litre with bleed valve.
393 | T- Piece PVC
394 | Cold Light (For placing vein) (with yellow/ white Bulb)) Should have light intensity controlled with smooth rotary potentiometer/ pressing button. 2) Should have output power 250 Watts (24 Volts)/ 150Watts (12 Volts). 3) Should have minimum dual control having 2 halogen/xenon/led lamps. 4) Should have SMPS based design ensures smooth working of light source within the voltage variation. 5) Should have fibre optic light cable 4.5mm - 10mm in diameter, 250cm300cm in length
395 | Transcutaneous Bilibinometer 1. Sample volume of < 100 μL required, automatic calibration facility 2. Total bilirubin concentration measurable (at least) in range of 0 to 30 mg/dl. 3. Time for total concentration measurement: ≤ 5 seconds 4. Should have filters: 455 and 575 nm (} 2%). 5. Should have error rate less than 5%. 6. Should have resolution- 0.1 mg/dl. 7. Automatic correction for Hemoglobin. 8. Measuring cell: Direct Hematocrit capillary readings. 9. heparinized hematocrit glass capillary.
396 | Infant Laser Temprature Gun 1.Measure mode: Body temperature and surface temperature - Measurement range:2. • Body temperature mode: 32 - 42.9℃ 3. Object temperature mode: 0 - 100℃ 4.Measuring accuracy(ambient temperature at 25℃) 0 - 31.9℃: ±1.5 ℃ 5. Sensing distance: 5 - 15cm 6.Display accuracy: 0.1 ℃ - Resolution: ±0.1 ℃
397 | OPHTHALMOLOGY Distance randot stereotest 4 targets (400 to 60 sec of arc) demonstration +C6:C22card and adult and Pediatric polarized glasses
398 | Teller acuity cards set of 17 cards with instructions for pediatric vision testing. easy to perform. 5 to 10 minutes to complete examination
399 | Tonopen fast, accurate, and reliable IOP measurements No Calibration Required. Less Sensitive to Corneal Centration. Gravity Independent
400 | Optokinetic Drum Black and White OKNBanner27x10inc. 8 inch size drum