## Sheet1
Sr.No | Item Name and Specifications Rate per Unit | Rate per Unit including all taxes | Items Manufacturer/ company name (compulsory) | Remark
551 | Microbiology Micropipette variable volume single channel, 0010 0100Microlitre
552 | Micropipette variable volume single channel, 0020 0200Microlitre
553 | Micropipette variable volume single channel, 0200-1000Microlitre
554 | Microslide cabinet Capacity of 10000 slides Horizontal, 75 x 25 mm SlidesThe usual size of the tray is 335 X 230 X35 mm
555 | Oil Cedar wood oil immersion oil 125 ml
556 | Oil immersion lens(100X) imported din type, antifungus treated, spring loaded for binocular microscope
557 | Disposable Sterile Petri dishes small size, 090mm Autoclavable Pack of 100
558 | Disposable Sterile Petri dishes small size, 150mm Autoclavable Pack of 050
559 | Slide Box (Plastic) Capacity 050 Slide
560 | Slide Box (Wooden) Capacity 050 slides
561 | Slide concave Blue star/Blue Ribbon/Gold, 1mm depth 75X25mm Size. , 01 Pack 50 slides,
562 | Slide plain Blue star/Blue Ribbon/Gold star size 75X25X1.2mm
563 | Stain CarbolFuschin Powder
564 | Stain Crystal violet Powder 100 gms
565 | Stain Crystal violet 125ml/bottle
566 | Stain Gentain violet Powder 100 gms
567 | Stain Gram Stain Kit 01 kit(125 ml of 1 bottle Gram, s Cystal violet + 125 ml of 1 bottle Gram, s decolorizer+ 125 ml of 1 bottle Safranin + 125 ml of 1 bottle Basic Fuschin )
568 | Stain Gram's Iodine Powder/Crystal 100 gms
569 | Stain Lugol's Iodine 125ml/bottle
570 | Stain Safranin Powder 100 gms
571 | Stain Safranin 125ml/bottle, 0.5% w/v
572 | Staining bottle rack plastic Height 3 inch, whole diameter 2 inch, capacity 8 bottle
573 | Staining bottle rack Wooden Height 3 inch, whole diameter 2 inch, capacity 8 bottle
574 | Sulphuric acid Concentrated (H2SO4)100 ml
575 | Swab Cotton bud with polypropylene stick ETO Steriledisposable,screw capped, Size 15cm×12mm diameter tube,
576 | Thermo meter Calibrated Minus 20 degree to 010 degree
577 | Thermo meter Digital calibrated Minus 000 degree to 250 degree Minus 000 degree to 250 degre
578 | Universal container Pack 50 nos, Sterile, ETO per piece, Capacity 50 60 ml, Polypropylene container, Diameter 5 6 cms, Height 5 7 cms
579 | Wire loop Nichrome double wound, Diameter 2.0mm (Volume 0.005ml)
580 | Xylene(sulphur free) 1X2.5 litre
581 | CarbolFucshin 125ml
582 | Phenol 500gm
583 | 25% H2So4 Solution 125 ml
584 | Methelene Blue solution 125 ml
585 | Gention Violate solution 125 ml
586 | Gram’s Iodine Solution 125 ml
587 | Safranine Solution 125 ml
588 | Acetone LR 2.5 Litre
589 | Lense Cleaning Solution 125 ml
590 | Oil Immersion 125 ml
591 | Concavity Slide pkt
592 | Oil Immersion Lense 125 ml
593 | Biosafety cabinet ClassII Type A2 (ESCO Global, Haier BioMedical, Thermo Fisher Scientific) General: Tender is invited from reputed original manufacture or their authorized agents/dealers only. Information regarding installation in India and satisfactory service and maintenance may be forwarded with all the details for verification. Technical: Certification : NSF 49/EN1249 or Equivalent standard Design : Approximately 4 Feet length X 2 feet Depth, Bio safety cabinets Class II, Type A2; 304 stainless steel interior Epoxy-coated steel exterior, Removable, seamless, dished work surface with lift out knobs Door- Fully closing, clear ¼” tempered safety glass sash Counter balanced with base stand. Circulation : Class 100, Supply and exhaust through HEPA filters. Inflow velocity of 105 fpm (0.5 m/sec), Down flow velocity of 55 fpm (0.3 m/sec), 70 % air recirculation Light : UV and sufficient illumination for work space. Gauges : For monitoring the condition of all HEPA filters as well as work space. Services Required : Installation & onsite validation, Calibration certificates Manuals : Operation, maintenance & part list with detailed specifications, Operational & maintenance Training Power Supply : Should include 210-240V/50 Hz Recommendation : Bio safety cabinets of Class II, Type A2 is recommended for all Viral Research Diagnostic labs (VRDL). Type B2 biosafety cabinets are recommended only for Biosafety Level 3 work
594 | Real time RT PCR (Applied Biosystems,Thermo Fisher Scientific, Eppendorf AG, Bio-Rad Laboratories, Inc.) General: Tender is invited from reputed original manufacture or their authorized agents/dealers only. Information regarding installation in India and satisfactory service and maintenance may be forwarded with all the details for verification. Technical: 1. Table top model. 2. Complete system including basic system, essential accessories, the state-of-art computer workstation, acquisition and analysis software, startup kit inclusive of calibration standards etc. 3. Open system to accommodate Taqman, SYBR green and all other fluorescent dye based chemistries. 4. Peltier based 96 well block 5. Standard optical 96 well plates, 0.2 ml strips, 0.2ml tubes compatibility 6. Min sample value requirement - 5µl 7. CCD camera with halogen/LED and at least five excitation and five emission filters 8. Multiplexing ability up-to five dyes in a single run 9. Calibrated dyes at installation: FAM/SYBR Green, VIC/JOE, NED/TAMRA/Cy3, ROX/Texas Red®, and Cy5, Should offer flexibility in dye selection. 10. Facility to calibrate new dye within the wavelength range without addition of new filters 11. Passive reference dye ROX or any other calibrated dye and should be optional 12. Option for melt curve analysis 13. Temperature range 4οC to 100οC 14. Sensitivity: Detection of 1 copy of template 15. Software applications: Comparative Ct, Standard Curve, Relative Standard Curve, Allelic Discrimination / SNP Genotyping, Plus/Minus, dissociation / melt curve 16. 220 V /50Hz. All accessories 17. CE mark or equivalent
595 | Refrigerator (450 liter Capacity) (Samsung India, LG India, Whirlpool) Technical specifications: Must be of a reputed brand, from manufacturer/authorized dealers having installation in reputed institutes/ firms, preferably in Maharashtra and service centers located in India. Necessary evidence to be provided. Design : Double door refrigerator, Inner chamber stainless steel 304 grade, Door locking system, magnetic door gasket, adjustable shelves , auto defrost facility. Capacity : 450 L Refrigerant : CFC Free Controls : moisture and humidity Alarm & Indicator : audio-visual alarm for door open and temperature fluctuation compressor on/off indicator Power Supply : Should include 210-240 V/50-60 Hz
596 | Centrifuge • Centrifuges are required in the Laboratory to separate various components of Blood for analysis. • Aerodynamic compact construction for vibration free performance • Table top version • Tube Capacity: No. 24 – 36 :Size 5 – 15 ml • Should have a digital timer • Body should be made of strong fabricated & corrosion resistant steel • Control panel – for start/stop switch, dynamic brakes, step less speed regulator with zero start switch & speed indicator with timer and protective fuses. • Door interlock. • Maintenance-free brushless drive motor with exact speed preselection and display. Speed range 100 to 6000 rpm and above, accuracy 1 rpm. • Centrifuge complete with Swig and basic rotors and four buckets- 01 set. • Tube Holders as appropriate • Shall meet IEC-60601-1-2 :2001(Or Equivalent BIS) General Requirements of Safety for Electromagnetic Compatibility. • The unit shall be capable of operating continuously in ambient temperature of 10 -40deg C and relative humidity of 15-90% • The unit shall be capable of being stored continuously in ambient temperature of 0 -50deg C and relative humidity of 15-90% • Power input to be 220-240VAC, 50Hz as appropriate fitted with Indian plug • Voltage corrector/stabilizer of appropriate ratings meeting ISI Specifications.( Input 160-260 V and output 220-240 V and 50 Hz) • The supplier should be ISO certified for quality standards. • Should be FDA, CE,UL or BIS approved product • Should comply with IEC/TR 61010-3-020: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 3-020: Conformity verification report for IEC 61010-2-020:1992 Particular requirements for laboratory centrifuges" • Comprehensive warranty for 2 years and 5 years AMC after warranty • User manual in English • Service manual in English • Certificate of calibration and inspection from factory • List of important spare parts and accessories with their part number and costing • Log book with instructions for daily, weekly, monthly and quarterly maintenance checklist. The job description of the hospital technician and company service engineer should be clearly spelt out.
597 | Elisa washer (Lisa Wash(Tulip Diagnostics Pvt. Ltd.), Biotek Instruments (I) Pvt Ltd.,Titertek) General: Tender is invited from reputed original manufacture or their authorized agents/dealers only. Information regarding installation in India and satisfactory service and maintenance may be forwarded with all the details for verification. Technical: Capability : 96 well microplates and strips, with flat, round, or “V” bottom well Manifold : 12 Channels Vacuum and fluid delivery : In built with Positive displacement syringe pump, drive Systems with adjustable flow rates Program : Memory for around 50 programs including dispensing Volume (50 to 300ul/well) and multispeed microplate shaking program Bottles : Wash, rinse and waste (volume 4-6 liter) Safety devices : Aerosol cover, removable Plate carrier, spill over protection and overflow protection safety system Display : LCD Display with Membrane Keypad Power Supply : 210-240V/50-60 Hz
598 | Elisa reader (Thermo Fisher Scientific, Biotek Instruments (I) Pvt Ltd., Bio-Rad Laboratories Inc., Tecan Inc) General: Tender is invited from reputed original manufacture or their authorized agents/dealers only. Information regarding installation in India and satisfactory service and maintenance may be forwarded with all the details for verification. Technical: Capability : 24, 48 or 96 well micro plates Spectral Range : 400-750nm, Accuracy +-1nm Absorbance : 0-4 O.D. Range Accuracy : Up to 0.001 O.D. Provision : Curve fittings formulas transformations & control assay validation, with compatible interface with PC & external printer. Accessories : Spare Lamps 2 Nos. Power Supply : 210-240V/50-60
599 | Laminar air flow (ESCO GLOBAL, Haier BioMedical, Thermo Fisher Scientific) General: Tender is invited from reputed original manufacture or their authorized agents/dealers only. Information regarding installation in India and satisfactory service and maintenance may be forwarded with all the details for verification. Technical: Certification : EN12469 OR equivalent Design : Vertical laminar flow cabinet, Approximately 4 Feet length X 2 feet depth, 304SS interior, epoxy/powder coated exterior Door fully closing with hinged sash. Support stand, solid one piece dished work surfaces Circulation : Down flow velocity of 40-70 fpm, ISO Class 5 or higher air Filters : HEPA filter, (99.99% efficient on particles 0.3 micron)/ULPA filter (efficiency of >99.999% at 0.1 to 0.3 micron sizes) provided with pre-filter) Light : UV and sufficient illumination for work space Gauges : For monitoring the condition of all HEPA filters as well as work space. Power Supply : Should include 210-240V/50-60 Hz
600 | Weighing balance (Mettler Pvt. Ltd., Sartorius AG, Contech India) General: Tender is invited from reputed original manufacture or their authorized agents/dealers only. Information regarding installation in India and satisfactory service and maintenance may be forwarded with all the details for verification. Technical: Single pan Analytical Balance with highest accuracy for weighing processes; readouts tohave at least four decimal places. Equipped with a draft shield chamber to eliminate interfering ambient effects. Weighing Range : 0.01 – 60 g Readability : 0.1 mg Calibration : External Display : LCD Display Verification interval : 0.001 g Pan Size : 80 - 100 mm Power Supply : 210-240V/50-60 Hz